## Slide 1: CMODIS and COCTS and its Radiation Calibration
20th WGCV meeting Hobart, Tasmania
Feb. 12-14, 2003
Xiaoxian Huang
Shanghai Institute of Technical Physics,Chinese Academy of Sciences

## Slide 2: Introduction
China had launched
One scientific research CMODIS on 25 March 2002 at Jiuquan launch center,
The first COCTS on 15 May 2002 at Taiyuan launch center

## Slide 3: Technical parameters of CMODIS
Channels | 34 | Cooler | Stirin(85K)
IFOV | 1.5mr | Accuracy | V-NIR-SWIR Better than 10%
Resolution | 500m (nadir) |  | TIR 1K
FOV | 45 | Quantization | 12bits
Scan rate | 20rpm | Scaning method | Two-surface mirror
Optical aperture | 110mm |  | 
Detector | Si HgCdTe 22 pixcel |  | 

## Slide 4: Band of CMODIS
Band | Channels
Visible(m) (0.403~0.803) | 20
Near IR(m) (0.823~1.023) | 10
Short Wave IR(m) (2.15~2.25) | 1
Thermal IR (m) (8.4~8.9)(10.3~11.4) (11.5~12.5) | 3

## Slide 5: Relative Spectral Response of CMODIS

## Slide 6: Technical parameters of COCTS
Channels | 10 | Cooler | Stirin(95K)
IFOV | 1.38mr | Accuracy | V-NIR-SWIR Better than 10%
Resolution | 1100m (nadir) |  | TIR 1K
FOV | 35 | Quantization | 10bits
Scan rate | 90rpm 45rpm | Scaning method | 45 rotate mirror K mirror
Optical aperture | 200mm |  | 
Detector | Si HgCdTe 4 pixcel |  | 

## Slide 7: Bands of COCTS
Visible and Near IR(m) 0.402~0.422,0.433~0.453 0.480~0.500,0.510~0.530 0.555~0.575,0.660~0.680 0.730~0.770,0.845~0.885 | Chlorophyll pollution atmosphere correction aerosol etc
Thermal IR (m) 10.30~11.40 11.40~12.50 | Temperature and ice

## Slide 8: Calibration of Visible and N-SWIR
Calibration of Visible and N-SWIR of CMODIS had been done under sunlight with diffuser panels in Dali Yunnan province China.
Calibration of Visible and N-SWIR of COCTS had been done using a Labsphere Integrating sphere(USS6500) in SITP.

## Slide 9: Calibration of Thermal IR
Calibration of thermal IR in both of CMODIS and COCTS had been done in KM-2 vacuum pipe.
For CMODIS the radiation from the blackbody on the CMODIS is the radiation reference.
For COCTS the cold space radiation is the reference and the blackbody on it is the on-orbit calibration source.

## Slide 10: In-site Calibration Works
After launch
The China National Satellite Ocean Application Service had spent much cost and efforts in testing the performance of CMODIS and COCTS in HangZhou gulf and Yellow-East Ocean of China.
The China National Satellite Ocean Application Service and National Meteorology Application Service have tested all remote sensors on Earth Observing Systems of China in Dunhuang and Qinghai Lake.

## Slide 11: Average field reflective ratio of Dunhuang

## Slide 12: Langley line of Dunhuang on June 4,2002 obtain by CE318 sun illuminant

## Slide 13: Image of Dunhuang obtained by CMODIS on June 4,2002SZ-3（R:15,G:20,B:3）

## Slide 14: Image of Qinhai lake obtained by CMODIS on June 4,2002SZ-3（R:15,G:20,B:3）

## Slide 15: Image before In-site calibration

## Slide 16: Image after In-site calibration

## Slide 17: Aerosol Optical Length Image over Bohai gulf obtained by CMODIS(June,11,2002)

## Slide 18: Conclusion
In-site calibration has high accuracy, but needs more time and cost.
We need to develop on-orbit calibrator especially in visible and NIR range in future remote sensor.
We are doing these works now at the FY-3, the second generation orbit Meteorology satellites of China.
We need to study more about calibration from another people who with such good experiments.

## Slide 19
I hope we have more opportunity to exchange with each other.
Thank you