## Slide 1
NIST Update -- CEOS WGCV 21
October 15-17, 2003
Beijing, P.R. China
Gerald T. Fraser
Optical Technology Division
Physics Laboratory
National Institute of Standards and Technology
http://physics.nist.gov/Divisions/Div844/div844.html
Notes: Thank you for this opportunity to brief you on NIST efforts in providing support for and direct calibration of remote sensing instruments.
The work I am going to talk about has been performed by my colleagues in the Optical Technology Division at NIST.
The Division has strong programs in radiometric physics and optical properties, and has obtained strong support from a number of government agencies, such as NOAA, DoD, and NASA, involved in satellite measurements.
I would like to begin by providing you an overview of some of our Division’s new facilities and capabilities , starting with a new laboratory.

## Slide 2
New Capabilities: HACR 2
Liquid
He
at 2K
Liquid
Nitrogen
Notes: HACR 2 is NIST’s second generation High Accuracy Cryogenic Radiometer being developed by Jeanne Houston and Joe Rice in our Division.
HACR is an electrical substitution radiometer which compares thermalized optical laser power to thermalized electrical power in a black cavity pictured below the schematic on the slide.
HACR 2 will become the primary standard at NIST for optical power responsivity, replacing HACR.
Expected improvements include
  A factor of two reduction in the power-measurement uncertainty to 0.01 % k = 1.
  A greater dynamic range, from 1 microwatt to 70 milliwatt, compared to the present 0.1 milliwatt to 0.8 milliwatt for HACR.
  Faster response time;
  Integration in the SIRCUS Laser facility to provide access to a near continuum of laser wavelengths from the infrared to the ultraviolet, improving the scale transfer to silicon trap detectors and other detectors.
  Module design allows new cavities to be developed for enhanced performance at specific wavelengths or powers.
  Greater liquid He hold time-- 9 days at K compared to the present 1 day at 5 K hold time for HACR.
The HACR 2 is presently undergoing testing and will be officially brought on line shortly after the move into the new Advanced Measurement Laboratory this winter. HACR 2 is NIST’s second generation High Accuracy Cryogenic Radiometer

## Slide 3
Tunable Laser
Wavemeter
Integrating
Sphere
Exit
Port
Spectrum
Analyser
Lens
Customer Detector
or Camera
to be Tested
Computer
Reference
Detector
Intensity
Stabilizer
Monitor Detector (output to stabilizer)
=325 nm - 18m
Chopper
or
Shutter
Speckle-removal
System
New Capabilities: SIRCUS
Notes: SIRCUS stands for Spectral Irradiance and Radiance Responsivity Calibrations using Uniform Sources facility developed by Keith Lykke, George Eppeldauer, Steve Brown, and Joe Rice at NIST
SIRCUS is a unique laser facility for making spectral irradiance and radiance responsivity measurements using
  Tunable, stable, high-intensity uniform laser radiation sources; and
  High-Accuracy transfer standard radiometers.
A schematic diagram of the SIRCUS facility is shown on the slide.
Briefly, it consists of an intensity-stabilized, single-frequency (known of 0.001 nm in the visible) tunable laser source (0.01 % power stability), a speckle removal system which destroys spatial coherence on the time and spatial scale of the measurements, an integrating sphere source (Spectralon coating, 2.5 cm diameter with an 8 mm diameter aperture or 30 cm diameter with a 7.5 cm output aperture or gold-coated sphere for the infrared) to provide a source of spectral radiance, and a reference detector, such a a silicon trap detector referenced against HACR or HACR 2.
SIRCUS comes in three colors:
  Ultraviolet-visible-near-infrared (325 – 1050 nm eventually extended to 2.5 microns)
  Midinfrared (1 to 5 microns, eventually extended to greater than 10 microns; CO and CO2 lasers)
  Traveling SIRCUS for visiting satellite and instrument calibration sites. (375 to 1050 nm)
Some features of SIRCUS include
  A high-power, uniform, monochromatic, and tunable source of radiance and irradiance
  Highest accuracy transfer and working standards
  The ability to calibrate spatially non-uniform detectors with low uncertainty in irradiance mode.
Consequently
  SIRCUS is the highest accuracy platform for detector-based calibrations
  The two SI units maintained by our Division, the candela and the kelvin above the silver melting point will soon be tied to the NIST HACR 2 through SIRCUS.
Uncertainties available from SIRCUS include
  Detector responsivity of 0.05 % for Si tunnel trap detectors
  Distance accuracy of 0.025 %
  Aperture areas specified to 0.02 % using a coordinate-measuring machine with an laser-interferometer measurement system.
  Uniformity of better than 0.1 %
  Exit port radiance of 0.1 %, correctable to 0.02 %
  Reference plane irradiance to 0.03 %.
Combined standard uncertainty of SIRCUS is 0.075 % k = 1; Inclusion of correlations will reduce this further.

## Slide 4
New Capabilities: Highly Accurate Absolute Radiometers
AP1
Notes: The development of the SIRCUS facility with its highly accurate spectral responsivity scale has motivated Howard Yoon, David Allen, and Robert Saunders at NIST to develop accurate and stable narrow-band absolute radiometers with absolute responsivity directly referenced to the HACR or the HACR 2.
The first of these radiometers is denoted AP1 for the first generation NIST Absolute Pyrometer.
The figure on the right shows a photograph of the AP1.
On the bottom is a schematic of the AP1 imaging a fixed point blackbody.
This radiometer was primarily developed to determine an improved value for the absolute freezing point of a pure gold fixed-point blackbody as part of an effort to improve the ITS-90 temperature scale, potentially for ITS-2010.
The AP1 weighs approximately 12 kilograms and is approximately 70 cm long including the cylindrical baffle.
Some of the features of the AP1 include
  set of three (12.5 mm) Invar optical rails for low thermal expansion and high rigidity.
  Gradient-index objective lens for low scatter and reflectance
  FOV (0.8 mm diameter target size) suitable for small aperture blackbodies and for comparisons with strip lamps
  Temperature controlled and stabilized filter-silicon-detector package. Detector (-15 C) and filter (30 C) independently temperature controlled.
  A narrow band (10 nm FWHM) interference filter centered at 650 nm with out of band blocking 10-6
  0.10 % uncertainty in radiance responsivity at 650 nm provided by SIRCUS, corresponding to an 80 millikelvin temperature uncertainty for the gold freezing point. 80 mK is the present difference between the ITS-90 and AP1 gold melting points which indicates that SIRCUS has met the stated uncertainties. Also, ITS-90 effectively has an uncertaintly as well, perhaps on the order of 40 mK. The SIRCUS radiance responsivity measurements are shown in the graph. Note the nearly 7 orders of magnitude in the dynamic range of the measurements.
  A size of source of 10-5, three orders of magnitude smaller than commercial instruments
It is anticipated that AP1 will lead to new absolute radiometers for the calibration of remote sensing instruments, photometers, and pyrometers with approximately 0.1 % uncertainty.
A push further to the blue, into the UV will lead to more accurate temperature measurements.
Absolute radiometers such as developed here will allow examination of higher temperature metal carbide eutectics, with melting points up to 3000 K or so, depending on composition.
These high temperature fixed points would further improve ITS-90 at the higher temperatures above the gold point and provide new high-tmperature fixed point blackbodies for radiometric standards.

## Slide 5
New Capabilities: FASCAL 2
Notes: FASCAL 2 is a new facility developed by Jim Proctor, Howard Yoon, and Charles Gibson to replace FASCAL for the calibration of FEL lamp spectral irradiance standards widely used for the calibration of sensors used by the remote sensing community.
FASCAL stands for Facility for Automated Spectroradiometric Calibrations.
The original FASCAL has been in continuous use since the early 1970’s.
Due to the increasing demands for spectroradiometric calibrations, particularly from the remote sensing community, NIST decided to build a new facility dedicated to spectral irradiance calibrations of 1 kilowatt FEL lamps.
The facility is designed to achieve a signal-to-noise ratio of greater than 1000:1 from 250 nm to 2500 nm with a bandwidth of 4 nm to 8 nm at a distance of 50 cm with a receiving aperture of 1 square centimeter.
The facility is uses an additive double Czerny-Turner double monochromator on a motorized stage with each component monochromator having dual gratings optimized for the ultraviolet and short-wave infrared wavelength regions.
Wavelength reproducibility is approximately 0.015 nanometers in the ultraviolet to near-infrared.
Available detectors include photomultipliers, 2-staged cooled Si, and extended-range InGaAs detectors.
The instrument has four lamp stations, a scale realization station using a high-temperature blackbody and filter radiometers, and a wavelength accuracy check station.
An FEL is automatically calibrated over the full range in less than 1.5 hours.
Consistent calibrations are obtained for FASCAL and FASCAL 2.
After further validation, all spectral irradiance FEL lamp calibrations will be shifted to FASCAL 2, most likely in the late winter or early spring of 2004.

## Slide 6
New Capabilities: LED Integrating Sphere Sources
Notes: Carlos Santana and Mark Lawrence, undergraduate students, Nordine Souaida, a visiting scientist from France, Steven Brown, Carol Johnson, Jim Proctor, and George Eppeldauer have developed a novel spectrally tunable integrating sphere light source based on light-emitting diodes for radiometric, photometric, colormetric, and display metrology applications.
The source can simulate natural sources, such as found in ocean color measurements, and can be used as a transfer source to propagate photometric and colormetric scales.
The initial prototype source used 40 LEDs (green, blue, and red) with 10 different spectral distributions. The next generation under development has approximately 150 LEDS arranged into 40 channels to better match target spectral distributions such a for ocean color.
The LEDs have a Gaussian spectral distribution with 20 nm widths and span the spectral range from 380 nm to 780 nm.
The current to each set of four diodes is individually controllable for spectral variation or matching, as judged by a spectroradiometer.
Future efforts will increase by a factor of four the number of distinct spectral sources. Also ultraviolet diodes offer new opportunties.
Simulations indicate that such a source could approximate standard light-source distributions over the 380 nm to 780 nm wavelength range to approximately 2 percent in luninance and 0.002 in chromaticity.
Funding has been procured for developing a blue source for use in calibrating the Marine Opitical Buoy or MOBY which is used for the vicarious calibration of ocean-color satellites such as SeaWiFS.

## Slide 7
Cryogenic Radiometry
Validation: Comparison of Detector-Based and
Source-Based Calibrations on a Set of Radiometers
ITS-90
SIRCUS: Detector Based
FASCAL: Source Based
Notes: Carol Johnson, Steve Brown, Keith Lykke, Charles Gibson and their NASA collaborators (Giulietta Fargion, Gerhard Meister, Stanford Hooker, Brian Markham, and Jim Butler) have compared two approaches for calibrating a set of four multichannel filter radiometers used in calibration efforts supporting satellite remote sensing programs:
  A source based scale referenced on the absolute scale provided by an electrical-substitution cryogenic radiometer; and
  A source based scale based on the ITS-90 temperature of the gold fixed-point blackbody.
Four radiometers
LXR – The LandSat Transfer Radiometer which is part of the Enhanced Thematic Mapper Plus (ETM+) calibration program.
SCR -- SeaWiFS Transfer Radiometer built to support calibration efforts in the SeaWiFS ociean-color satellite program.
SXR-II -- Used for source comparisons in the Sensor Intercomparison and Merger for Biological and Interdisciplinary Ocean Studies (SIMBIOS) program.
VXR – Visible Transfer Radiometer is used in EOS calibration comparisons.
were calibrated for absolute radiance responsivity, r(lambda) in the figure above, using the a laser-illuminated integrating sphere source with absolute scale referenced to a electrical-substitution cryogenic radiometer. The calibrations were performed using the SIRCUS facility.
The NIST Facility for Automated Spectroradiometric Calibrations (FASCAL), the 2nd generation of which was described previously, was used to calibrate a lamp-illuminated integrating sphere source for spectral radiance, L(lambda) in the figure above, from 300 nm to 1100 nm by comparison to a variable temperature blackbody referenced to a gold freezing-point blackbody. The integrating sphere source is the EOS/NIST portable radiance source (NPR), a 30 cm diameter, internally illuminated, Spectralon-coated sphere.
The resulting values for r(lambda) and L(lambda) were used to predict the signal for the radiometers, SP, when viewing the integrating sphere source.
The radiometers also directly examined the NPR integrating sphere source and the signal SM were determined. As shown in the figure where we plot the ratio of SP/SM against wavelength, the ratio of SP/SM is nearly unity. Here in the figure the legend identifies the plotting symbols. The lamp configuration for the integrating sphere is indicated in parentheses.
Indicating that the results generally agree within the combined uncertainties (shown for a few cases), mutually validating the two independent approaches.

## Slide 8
Validation: Comparison of Laser-Based and Conventional
Calibration of Sun Photometers
Notes: Sun photometers, such as the Simbad A set of 11 filter radiometers pictured, are used to characterize the radiative properties of the atmosphere and of the oceans—so called “sea radiances” for the SeaBASS program in support of ocean-color satellite validation.. For atmospheric studies, they measure both the incident solar irradiance as well as the sky radiance (from scattered incident flux). Global networks of sun photometers provide data products such as aerosol optical thickness derived from these measurements.
In satellite-based ocean color remote sensing measurements, approximately 90 % of the signal measured at the satellite sensor comes from atmospheric scattering. Proper atmospheric characterization and aerosol modeling is necessary to separate the atmospheric component from the top-ofthe-atmosphere radiance, required for valid satellite-based measurements of ocean color data products such as global, near-surface, phytoplankton chlorophyll-a distributions.
Sun photometers are usually calibrated for irradiance and radiance responsivity—with the former being done by the standard Langley-Bouguer cross-calibration technique against a primary reference sun photometer as shown schematically on the right while the later is obtained using a standard integrating sphere source as shown in the lower left.
Nordine Souiaidia, a guest scientist, collaborated with NIST colleagues Steve Brown, Carol Johnson, and Keith Lykke, and NASA SIMBIOS Project and Scripps Institute scientist to compare these two approaches with that obtained using a integrating sphere pumped by a single-frequency laser system—SIRCUS.
They found for Top of the Atmosphere spectral irradiances derived from the Langley plot that good agreement is obtained for the regions where the solar irradiance spectrum is slowly varying.
A 0.15% difference was found for the 750 nm channel and 5% was found for the 440 nm channel
The SIRCUS measurements also revealed center frequency shifts for most of the channels, up to 2 nm in some case, indicating the importance of careful characterization.
SIRCUS calibrations for radiance and irradiance of sun photometers is better than 0.1 %, indicating the potential for better characterization of the atmospheric column through comparison of exo-atmospheric measurements with well calculated sun photometer measurements.

## Slide 9
Intercomparisons: Aperture Measurements for Solar Irradiance
Notes: Knowledge of the aperture area is critical in the conversion of the measured radiant flux to solar irradiance.
To assess the accuracy of aperture measurements used in exo-atmospheric solar-irradiance measurements, an Earth Observing System (EOS)-sponsored international comparison of aperture area measurements of limiting apertures provided by solar irradiance researchers is under way, lead by Toni Litorja and Carol Johnson of NIST and Jim Butler of the EOS Project Science Office.
This measurement comparison is being made in response to issues address at the May 2000 calibration workshop for the Total Irradiance Monitor (TIM) Instrument on the Earth Observing System (EOS) Solar Radiation and Climate Experiment (SORCE). Workshop participants endorsed the idea of holding a comparison program on aperture area measurement with NIST serving as the lead institution.
The workshop participants realized that exo-atmospheric solar-irradiance measurements made over the past 25 years incorporated limiting apertures measured by a number of laboratories using a variety of techniques.
Exo-atmospheric solar irradiance measurements show instrument-to-instrument variation on the order of 7.5 W/m2 out of a total irradiance of approximately 1366 W/m2, or about 0.55 %. This difference is within the 0.2 % to 0.3 % k = 1 uncertainty of the irradiance measurements, however a close examination of the measurements suggest the presence of one or more systematic causes. Aperture area issues could be one possibility. Ideally, the aperture area would be known to something better than 0.1 %, implying that the aperture radius must be known to better than 0.05 %. For a 2.5 mm radius aperture this corresponds to 1.25 microns absolute distance uncertainty. Additional concerns not discussed here is change in the aperture area from chemical effects, such as oxidation before launch.
Apertures were measured at NIST using the Absolute Aperture Measurement Facility at NIST pictured in the lower left with a schematic in the upper left.
As shown in the schematic, the measurement technique employs a coordinate measurement machine which uses non-contact video microscopy for edge detection, with the aperture mounted on high-precision laser-interferometer stage for coordinate measurements.
The aperture area comparison aims to quantify the relative differences between the participating institutions’ aperture area measurements. Apertures used in solar irradiance measurements have a variety of shapes due to variation in fabrication methods and instrument design philosophy. Possible shapes are shown in the lower right of the slide.
To date, apertures have been received from two participating institutions, referred to as A and B in the figure above. It is anticipated that apertures will be received from a third and possibly a fourth institution. The results presented here should thus be taken as preliminary.
The preliminary results are shown in the graph. Six apertures were measured for each of the two participating laboratories. The area, as reported by the laboratory, is normalized by the area determined by NIST and plotted as a function of aperture designation or number. The vertical lines correspond to the combined uncertainty in this ratio (k = 2), based on the uncertainties in area as reported by the participants and the uncertainties for the NIST measurements.
The present measurements show that at least for laboratories A & B, aperture variation account for less than 1 W/m2 in solar irradiance value, significantly less than the 7.5 W/m2 variation seen between instruments.

## Slide 10
Intercomparisons: SIMBIOS Ocean Color Radiometric Scales
The Second SIMBIOS Radiometric Intercomparison (SIMRIC-2), March – November 2002,
Meister et al., NASA/TN-2002-210006, Vol. 2, Aug. 2003
Laboratory | Primary Calibration Source
NRL Optical Sensing Section | FEL
NASA Code 920.1 Calibration Facility | FEL NIST
Wallops Flight Facility (NASA) | Non FEL Irradiance Standards
Moss Landing Marine Laboratory | Integrating Sphere NIST
Scripps | Integrating Sphere
Biospherical Instruments | FEL NIST
UCSB | FEL NIST
U. South Florida | Integrating Sphere
University of Miami | FEL
Satlantic | FEL
Notes: NIST was indirectly involved in the Second SIMBIOS Radiometric Intercomparison (SIMRIC-2) held between March to November 2002 and led by Gerhard Meister of the SIMBIOS Project.
SIMBIOS Funding is now cancelled, but was an excellent example of a program with strong international linkages.
SIMBIOS refers to Sensor Intercomparison and Merger for Biological and Interdisciplinary Oceanic Studies (SIMBIOS) Project
The purpose of SIMRIC is to ensure consistency and document approaches in establishing the spectral radiance scales for laboratories involved in ocean-color related measurements.
The NIST-designed, Reyer Corporation built, SeaWiFS Transfer Radiometer (SXR-II) measured the standard radiances produced by a serious of ten participating laboratories listed above at six wavelengths from 411 nm to 777 nm.
The SXR-II measured radiances were compared with the values claimed by the laboratories.
The SXR-II was calibrated three times at NIST during the study, using the SIRCUS facility described previously.
The various laboratories used integrating spheres or FEL-illuminated Spectralon plaques to obtain a standard source of spectral radiance.
A number of the FEL lamps and one integrating sphere source were calibrated for spectral irradiance (FEL) or spectral radiance (ISS) at NIST.
The measurement results showed agreement within the combined uncertainties for all but two laboratories [Miami and Wallop]. (2% uncertainties for lab scales, 1% uncertainty in SXR-2 measured radiance)
[issues for Wallop-- wrong reference point on the lamp for establishing the distance scale
Issues for Miami—Spectralon plaque not calibrated; a reflectance of 0.98 was assumed.]
The authors report that measured radiances are typically greater than expected radiances, suggesting a stray light issue. This same result was observed in SIMRIC-1
In general the study verified the approaches used to establish the radiometric scales for calibrating the SeaBASS “ground” truth data used for assessing the quality of satellite ocean-color measurements such as provided by SeaWiFS.

## Slide 11
Instrument Calibration: Robotic Lunar Observatory (ROLO)
Notes: Goal is to establish the moon as an absolute standard of spectral radiance for the calibration of remote sensing satellites in the reflected solar region (to the blue of 2.5 microns).
The advantage of using ROLO as a standard source of spectral radiance:
Surface reflectance is stable to better than 1 part in 10 million.
The moons brightness is within the dynamic range of most nadir viewing imaging satellites.
In principle, the moons radiance can be known to better than a percent.
Observation of the moon is relatively straightforward for most satellite sensors.
Two calibration approaches are currently used. One a pure laboratory calibration, the other is a stellar calibration.
Laboratory calibration:
A NIST calibrated FEL irradiance-standard lamp illuminates a Spectralon plaque to create a standard source of spectral radiance, which is periodically verified by ROLO. The radiance is known to 2 %.
This approach has problems since the radiance source is Lambertian and in the near field while the ROLO telescopes are focused at infinity.
Exo-atmospheric calibration using the star Vega:
The absolute spectral irradiance is effectively based on photometric measurements by astronomers.
The analysis must correct for the atmosphere.
A four-year analysis of ROLO day yields the exo-atmospheric irradiance of VEGA to better than 1 percent.
Comparison of the calibration by Vega and lamp-plaque approach give a 20 to 30 % discrepancy.
This discrepancy may not be surprising since the full angle viewed by the moon is 0.5 degrees, by VEGA near 0 degrees, and by the lamp-plaque nearly 8 degrees.
NIST calibration scientists Steve Brown and Carol Johnson and NASA calibration scientist Jim Butler have been working with the ROLO team on trying to provide an absolute scale for ROLO observations of the moon.
A collimator has been developed to mimic the radiation field of the moon as seen by ROLO.
The collimator is a spherical mirror illuminated on axis at the focus by a source of known spectral radiance. The radiance of the source known from FASCAL and the mirror reflectance (NIST) give the known radiance for ROLO.
Preliminary results indicate spectrally varying discrepancy between the collimator-based and Vega-based calibration.
This results is puzzling since ROLO and satellite measurements of the lunar irradiance don’t reveal such a discrepancy.
Future efforts are targeted at resolving this discrepancy—one possibility is the intervening atmosphere between VEGA and ROLO not present between the plaque and ROLO.

## Slide 12
Instrument Calibration: Marine Optical Buoy (MOBY)
Notes: My colleagues have discussed this effort previously at the working group, but I would like to provide you an update of our activities.
As a reminder, MOBY provides “vicarous” calibration of ocean-color satellites such as SeaWifs and MODIS. A picture of the buoy is shown on the right, alongside a minivan. The buoy when in is deployed must have its windows routinely cleaned by scuba divers.
Unfortunately for Carol Johnson and Steve Brown of our Division, the project leads, the buoy is floating off the coast of Hawaii, requiring frequent, or overly frequent, trips to the islands.
In fact, it is not a vacation, requiring work in the open air and in tents pictured above.
Instruments involved in ocean color measurements are typically calibrated against incandescent sources such as lamp-illuminated integrating spheres or reflectance plaques. These types of calibration sources have very different spectral distributions than the distribution of up-welling or water-leaving radiance. Because of the differences in the spectral distributions of the calibration source and the ocean color, stray light (or spectral out-of-band light) can introduce significant error in the measurement.
Direct evidence for such stray lights effects were seen in the characterization of MOBY using our Traveling SIRCUS system, consisting of a calibrated tunable-laser illuminated integrating sphere source.
The MOBY sensor, called the Marine Optical System (MOS), is located in the bottom of the buoy, and is comprised of two CCD-based spectrographs; one covers the 350 nm to 640 nm spectral range (called the Blue Spectrograph (BSG)), and the second covers the 550 nm to 940 nm spectral range (called the Red Spectrograph (RSG)). The MOS analyzes light propagated down optical fiber from sensors shown on the arms of the buoy.
Characterization of the dual spectrograph system with single-color laser sources reveal significant stray-light effects. The full effect of stray light on ocean carbon level estimates await a detailed analysis of all three components associated with the instrument calibraiton: the satellite, MOBY, and field-ocean radiometers.
Additional work on MOBY includes the development of a robotic light collection system fiber coupled to the MOS spectrograph which no longer needs to be sitting at the end of the buoy. The robotic system allows instrument shadows from the buoy to be avoided. Such shadow affect the measurements particularly in turbid waters.
Work is also in progress to reprocess the MOBY data set, considering, for instance, corrections for stray light.

## Slide 13
Workshop: Satellite Instrument Calibration for
Measuring Global Climate Change
Sponsors: NIST, NPOESS-IPO, NOAA, and NASA
When: November, 2002
Where: University of Maryland
Organizing Committee: Raju Datla, Chair, NIST
  Mike Weinreb, NOAA
  George Ohring, Consultant to NOAA
  Steve Mango, NPOESS-IPO
  Jim Butler, NASA
  Dave Pollack, UAH
  Purpose: “Develop requirements and recommend directions for future improvements in satellite instrument characterization, calibration, inter-calibration, and associated activities, to enable measurements of global climate change.”
Notes: Read from slide.
Approximately 75 attendees—actually participants
Workshop focus:
  Passive satellite sensors - ultraviolet to microwave
  Climate variables
    Solar irradiance, Earth radiation budget, and clouds
      Total solar irradiance, spectral solar irradiance, outgoing longwave radiation, net incoming solar radiation, cloudiness
    Atmospheric
      Temperature, water vapor, ozone, aerosols, precipitation, and carbon dioxide
    Surface
      Vegetation, snow cover, sea ice, sea surface temperature, ocean color
Recommendation or themes include--
  Elevate climate calibration requirements to critical importance in NPOESS
  Develop characterization requirements for all instruments and insure that these are met
  Conduct and verify pre-launch calibration of NPOESS and GOES-R instruments using NIST transfer radiometers
  Simplify the design of climate monitoring instruments
  Implement redundant calibration systems
  Establish means to monitor the stability of the satellite sensors
Finally, the organizers are considering proposing a follow-on workshop to discuss implementation issues.

## Slide 14
Strategic Plan for the U.S. Climate Change Science Program
Objectives to help Observing and Monitoring System Goal 3: Provide Stewardship of the Observing System
Objective 3.1. Follow climate monitoring principles. … Instrument calibration, characterization, and stability become paramount considerations. Instruments must be tied to national and international standards such as those provided by the National Institute of Standards and Technology (NIST). … Agreed-upon measurement protocols and procedures (e.g., continuous data validation and intercomparisons) are required to produce climate quality data.
Objective 3.2. Provide independent measurement and analysis. … Experience from previous scientific assessments, the NIST, and other national standards laboratories have shown that actual accuracy is only known after comparison of independent measurements and analysis from multiple laboratories. …

## Slide 15
New Facilities: The New Advanced Measurement Laboratory
Notes: Many of the Division Programs in support of satellite remote sensing will be moving into the $235 million Advanced Measurement Laboratory.
Occupation is scheduled to start in December 2003 and be complete by June 2004.
SIRCUS, HACR 2, and the Aperture Measurement Instrument are scheduled to move into the new laboratory.
The AML will provide
  Significantly improved air quality to 35 particles/liter or less.
  Baseline temperature control to 0.25 degrees Celsius.
  Reduced vibration control to better than 3 micrometers/second.
  Humidity control to 5 %.
  Improved electrical power supply conditioning.

## Slide 16
Acknowledgements
NIST Optical Technology Division
Collaborators
Leaders of the NIST Calibration Effort
  Carol Johnson
  Joe Rice
  Steve Brown
Other NIST Collaborators
  David Allen
  Raju Datla
  Ted Early
  Charles Gibson
  Toni Littorja
  Keith Lykke
  Al Parr (Division Chief)
  Jim Proctor
  Bob Saunders
  Howard Yoon
SeaWiFS
Numerous Other Collaborators, Including
  Jim Butler, NASA Goddard
  Dennis Clark, NOAA
  Hugh Kieffer, USGS
  Chuck McClain, NASA Goddard
Notes: Thank you again for this opportunity to present here.
I want to acknowledge my colleagues, particularly Carol Johnson, Steve Brown, and Joe Rice, who have been leading our remote sensing calibration effort.
Joe, in particular, has been active in connecting us with the NPOESS
Strong support and collaboration has been provided by NOAA, NASA and their EOS office, the SeaWiFS program, and the Air Force.