## Slide 1
NIST Agency Report -- CEOS WGCV 23
March 7 – 10, 2005
Córdoba, Argentina
Robert D. Saunders
DSK Science, Inc.
Optical Technology Division
Physics Laboratory
National Institute of Standards and Technology
http://physics.nist.gov/Divisions/Div844/div844.html
Notes: Thank you for this opportunity to brief you on NIST efforts in providing support for and direct calibration of remote sensing instruments.
The work I am going to talk about has been performed by my colleagues in the Optical Technology Division at NIST.
The Division has strong programs in radiometric physics and optical properties, and has obtained strong support from a number of government agencies, such as NOAA, DoD, and NASA, involved in satellite measurements.
I would like to begin by providing you an overview of some of our Division’s new facilities and capabilities , starting with a new laboratory.

## Slide 2: Outline
What’s new with the Core Facilities
Accomplishments, Interactions, and updates in Earth observing, Remote Sensing Projects

## Slide 3
Absolute Radiometry
HACR 2
Electrical substitution radiometry
Optical watt equivalent to electrical watt
Cryogenic cavities and flux-stabilized lasers
Reference standard detectors—flux responsivity
Scale realizations in detector responsivity, spectro-radiometry, and photometry.
HACR 2
Modular & horizontal cavity, windowless operation
Testing and characterization completed
L-1 ACR
Spectral responsivity scale realization completed; CCPR-K2.c spectral responsivity (200 nm to 400 nm) in progress
L-1 ACR
Notes: Joe has spent the last several months bringing the HACR2 online. I think this is with a new cavity, after evaluating the results from the first cavity. So I say in the bullets “testing and characterization completed” Likewise, Steve and Keith have gotten L-1 operational and performed characterizations etc. Best to talk to Joe and or Steve for more details.
Notes from WGCV 22 slide:
HACR 2 is NIST’s second generation High Accuracy Cryogenic Radiometer being developed by Jeanne Houston and Joe Rice in our Division.
HACR is an electrical substitution radiometer which compares thermalized optical laser power to thermalized electrical power in a black cavity pictured below the schematic on the slide.
HACR 2 will become the primary standard at NIST for optical power responsivity, replacing HACR.
Expected improvements include
  A factor of two reduction in the power-measurement uncertainty to 0.01 % k = 1.
  A greater dynamic range, from 1 microwatt to 70 milliwatt, compared to the present 0.1 milliwatt to 0.8 milliwatt for HACR.
  Faster response time;
  Integration in the SIRCUS Laser facility to provide access to a near continuum of laser wavelengths from the infrared to the ultraviolet, improving the scale transfer to silicon trap detectors and other detectors.
  Module design allows new cavities to be developed for enhanced performance at specific wavelengths or powers.
  Greater liquid He hold time-- 9 days at K compared to the present 1 day at 5 K hold time for HACR.

## Slide 4
SIRCUS
Detector – based facility for determination of absolute spectral irradiance or radiance responsivity of optical sensors
Reference detectors calibrated using cryogenic electrical substitution radiometry; laser illuminated integrating sphere sources for monochromatic characterization and calibration of radiometers; variety of laser systems provide tunable coverage; low uncertainties: 0.1 % to 0.2 % (k = 2)
Continued research effort for stray light correction in single grating CCD or PDA commercial spectrographs
Spectral Irradiance and Radiance responsivity Calibrations using Uniform Sources (SIRCUS)
Notes: SIRCUS stands for Spectral Irradiance and Radiance Responsivity Calibrations using Uniform Sources facility developed by Keith Lykke, George Eppeldauer, Steve Brown, and Joe Rice at NIST
SIRCUS is a unique laser facility for making spectral irradiance and radiance responsivity measurements using
  Tunable, stable, high-intensity uniform laser radiation sources; and
  High-Accuracy transfer standard radiometers.
A schematic diagram of the SIRCUS facility is shown on the slide.
Briefly, it consists of an intensity-stabilized, single-frequency (known of 0.001 nm in the visible) tunable laser source (0.01 % power stability), a speckle removal system which destroys spatial coherence on the time and spatial scale of the measurements, an integrating sphere source (Spectralon coating, 2.5 cm diameter with an 8 mm diameter aperture or 30 cm diameter with a 7.5 cm output aperture or gold-coated sphere for the infrared) to provide a source of spectral radiance, and a reference detector, such a a silicon trap detector referenced against HACR or HACR 2.
SIRCUS comes in three colors:
  Ultraviolet-visible-near-infrared (325 – 1050 nm eventually extended to 2.5 microns)
  Midinfrared (1 to 5 microns, eventually extended to greater than 10 microns; CO and CO2 lasers)
  Traveling SIRCUS for visiting satellite and instrument calibration sites. (375 to 1050 nm)
Some features of SIRCUS include
  A high-power, uniform, monochromatic, and tunable source of radiance and irradiance
  Highest accuracy transfer and working standards
  The ability to calibrate spatially non-uniform detectors with low uncertainty in irradiance mode.
Consequently
  SIRCUS is the highest accuracy platform for detector-based calibrations
  The two SI units maintained by our Division, the candela and the kelvin above the silver melting point will soon be tied to the NIST HACR 2 through SIRCUS.
Uncertainties available from SIRCUS include
  Detector responsivity of 0.05 % for Si tunnel trap detectors
  Distance accuracy of 0.025 %
  Aperture areas specified to 0.02 % using a coordinate-measuring machine with an laser-interferometer measurement system.
  Uniformity of better than 0.1 %
  Exit port radiance of 0.1 %, correctable to 0.02 %
  Reference plane irradiance to 0.03 %.
Combined standard uncertainty of SIRCUS is 0.075 % k = 1; Inclusion of correlations will reduce this further.

## Slide 5
NIST Spectral Irradiance Scale Uncertainties
CCPR-K1.a Spectral Irradiance Comparison (1000 W tungsten quartz halogen lamps, 250 nm to 2500 nm)
k = 1
Notes: These are the reported uncertainties by the participants in the CCPR irradiance comparison. The NIST values are the lowest. NIST uses a detector based technique for spectral irradiance, implemented in 2000. These are out of the Draft A report. The Draft B report has been approved and btw Howard says there are no major changes to this plot.
The key comparison reference value is a weighted mean. The reported uncertainties were used to weight the data, EXCEPT if they were below the cut off value (solid line). In that case, the solid line was used to weight the data (this is what happened to the NIST results, especially evident in the SWIR. The NIST results agree with this weighted mean to within k=2 for the entire wavelength range. If the KC ref. value were calculated using the NIST reported uncertainties, not the cut off values, our results would be even closer to the KC ref. value.

## Slide 6
NIST UV Spectral Irradiance Scale
Comparison of the hydrogen arc FASCAL scale to the new SURF III scale.
5x reduction in irradiance uncertainty
SURF III irradiance calculable - Schwinger Eqn
Improvements in lamp reproducibility
Elimination of transfer based on relative spectral irradiance calibration using the hydrogen arc
CCPR-K1.b participation (200 to 400 nm)
Impact: UV radiometry, ozone retrievals

## Slide 7
Detector – based Kelvin
Absolute Pyrometer (AP1)
Absolute pyrometers developed and calibrated on SIRCUS
Eutectic blackbodies being developed for high-T standards
Absolute pyrometer and eutectic standards validated by intercomparison with NPL

## Slide 8: Outline
What’s new with the Core Facilities
Accomplishments, Interactions, and updates in Earth observing, Remote Sensing Projects

## Slide 9
Aperture Measurements for Solar Irradiance
In support of cal/val for EOS, heritage apertures are being measured for absolute area
Measurements for a third lab are in progress
TSI Uncertainty workshop scheduled for July 2005 at NIST
NIST Aperture Area Measurement Facility
Total Solar Irradiance
Notes: http://spot/colorado.edu/~koppg/TSI/

## Slide 10
Novel Sources for Remote Sensing
Solid-state LED Calibration Sources
Advantages:
  Controllable output 350 nm to >900 nm;
  Higher flux in UV and blue;
  Compact, light-weight, low power consumption;
  Good temporal stability;
  Shock resistant;
  Provides a great deal of flexibility in source design;
  Enables derivation of sources with unique spectral distributions.
Field Source with MOBY radiance port
Calibration Source schematic for MOBY
Notes: We have made and deployed a small LED based field source for NOAA “small boat” operations to check their robotic, fiber-optic coupled spectrograph. We are modifying the coupling optics and electronics.
We are also making a calibration source for the Marine Optical Buoy (MOBY) project that can replace the TQH lamp –based integrating sphere sources. The idea is to match the blue (Case 1) waters and get better uncertainties because the SNR in the blue will be less (more calibration flux). The sphere (schematic shown) will be illuminated by several types (colors) of LEDs.
We are also working on sources that have tunable, (programmable) spectral distributions. This source can simulate natural sources, such as found in ocean color measurements, and can be used as a transfer source to propagate photometric and colorimetric scales.

## Slide 11
New Remote Sensing Laboratory
Cleanroom Facilities
16 m x 7 m Laboratory
Environmental Chamber
Fume Hood
Traveling SIRCUS
Spectral Irradiance & Radiance Facility
Shipping & Storage
Electronics Bench
1 m Sphere
Optical-Mechanical Components
Notes: We are still moving in, but will allow for unified projects (LED source development, transfer of FASCAL radiance scales to much lower or higher or different shapes specific to remote sensing, testing and validation of instruments, etc.

## Slide 12
Project Summary
Earth Observing System (EOS)
Jim Butler, NASA/GSFC cal/val lead
Primary efforts, FY05: TSI workshop, calibration of Robotic Lunar Observatory telescopes
National Polar Orbiting Environmental Satellite System (NPOESS) and NPOESS Preliminary Project (NPP)
Karen St. Germain and Steve Mango, IPO
Primary efforts, FY05: OMPS cal/val at Ball, CrIS blackbody at NIST with TXR, publication of TXR verification of SBRS VIIRS blackbody radiance
Geostationary Operational Environmental Satellite (GOES) – R
Mike Weinreb, NOAA/NESDIS
Primary efforts, FY05: Plan for ABI calibration verification efforts; application of TXR measurements of the GOES Imager blackbody source
NIST continues to collaborate with Earth observing programs to assess the accuracy of the radiometric characterization and calibration of flight sensors, as well as field equipment.
Notes: Carol is EOS and GOES-R; Joe Rice is NPP/NPOESS
OMPS Ozone Mapping and Profile Spectrometer (close enough)
Cross-track Infrared Sounder (CrIS)
ABI advanced baseline imager

## Slide 13
Recommendations to Plenary
Artifacts for down-stream characterization (programs should produce and archive “witness samples”)
“Witness samples” could be from the flight set, in order to ensure reproducibility of determined parameters.
Case in point: filters for sounding channels (C. Cao).

## Slide 14
Acknowledgements
NIST Optical Technology Division
Collaborators
Leaders of the NIST Calibration Effort
  Carol Johnson
  Joe Rice
  Steve Brown
Other NIST Collaborators
  David Allen
  Raju Datla
  Charles Gibson
  Toni Littorja
  Keith Lykke
  Al Parr (Division Chief)
  Jim Proctor
  Bob Saunders
  Howard Yoon
SeaWiFS
Numerous Other Collaborators, Including
  Jim Butler, NASA Goddard
  Steve Mango, IPO
  Dennis Clark, NOAA
  Steve Kirkner, NOAA
  Tom Stone, USGS
  Chuck McClain, NASA Goddard
Notes: Thank you again for this opportunity to present here.
I want to acknowledge my colleagues, particularly Carol Johnson, Steve Brown, and Joe Rice, who have been leading our remote sensing calibration effort.
Joe, in particular, has been active in connecting us with the NPOESS
Strong support and collaboration has been provided by NOAA, NASA and their EOS office, the SeaWiFS program, and the Air Force.