## Slide 1
Field Comparison Considerations
CEOS/WGCV/IVOS
March 2 – 11, 2005
Córdoba, Argentina
Robert D. Saunders
DSK Science, Inc.
Optical Technology Division
Physics Laboratory
National Institute of Standards and Technology
http://physics.nist.gov/Divisions/Div844/div844.html

## Slide 2: Outline
Heritage
Ideas & suggestions for future comparisons

## Slide 3
Heritage
HACR 2
L-1 ACR
Example Activities
North American Interagency Intercomparison of Ultraviolet Monitoring Spectroradiometers Purpose
  Assess accuracy of downwelling surface irradiance measurements in the UV;
  Perform limited field characterization and calibration using NIST standards
SeaWiFS Intercalibration Round-Robin Experiments (SIRREXs)
Purpose
  Ensure that radiometric standards used by institutions involved in the validation of SeaWIFS measurements were consistently calibrated with the values traceable to stated reference standards (e.g., SI);
  Clear articulation, demonstration, and training related to the community-accepted Ocean Optics Protocols
1997 Lunar Lake, Nevada Surface Reflectance to Top of the Atmosphere Radiance
Purpose
  Laboratory characterization of field radiometers
Notes: We have publications if they want them they should email us.

## Slide 4
Ideas & Suggestions (1)
Common Broadband Source for Intercomparison
Whereas
  Field radiometers are used outdoors and in remote locations, which may result in operation on 1) battery power; 2) very short warm-up times; and 3) extreme values for ambient (air) temperature and relative humidity
and
  Solar-illuminated reflectance standards are not usually viewed simultaneously, so there is no way to intercompare radiometers or assess their radiometric stability during field exercises
it follows that the comparison results can be no better than those obtained in a laboratory under more ideal conditions.
We propose to assess this uncertainty component by pre (or pre and post) laboratory comparisons.
Technique
  Stable source (e.g. integrating sphere with monitor detectors and/or external NIST monitor radiometer (the source does not have to be calibrated);
  Field radiometers measure the source under controlled conditions

## Slide 5
Example result: short term stability
Figure 5, Brown et al. 2001, Remote Sensing of Environment 77, 367-376
Difference in sphere radiance measurements using the UA Fieldspec FR. The total elapsed time between measurements was 5 min (solid line), 25 min (open squares) and 45 min (open triangles.

## Slide 6
Related Tests
Common Broadband Source – other uses
Battery vs. ac power
Cold start vs. fully warmed up instrument
Field-of-view tests: independent of spectrograph sub-system? (radiance measurements should be independent of distance for a uniform, extended source)
Sensitivity to fiber removal and insertion; position of optical fiber
Sensitivity to alignment & realignment
Linearity
Integration time
NIST Portable Radiance source (NPR)

## Slide 7
Ideas & Suggestions (2)
Standard sources in the field
This is more difficult at remote locations typically used for vicarious calibration, but may be possible for a dedicated intercomparison activity
Necessary elements:
  Engineered for extreme stability (including sensitivity to temperature)
  Tested thoroughly, e.g. in an environmental chamber
  Monitor photodiodes or spectroradiometer
  Automated data acquisition of relevant parameters
  Control of stray radiation (including sunlight)
Field Irradiance Calibration
←Old way, had to run at night
New way, light-tight, run anytime→
Notes:   We have also done this with LED sources.

## Slide 8
Ideas & Suggestions (3)
Novel sources
Spectral out of band in array spectrographs is a severe problem when the calibration source and the Earth scene (measured source) differ in relative spectral shape
Specular
Haze
Diffuse
Incident Ray
Surface
Imperfect gratings lead to non-delta function spatial images in monochromatic light
Resulting spectral biases contribute unequally, depending on spectral shape of the source

## Slide 9
Evaluate for Presence of Effect (1)
Novel sources with well defined and distinct spectral shapes (compared to tungsten quartz halogen sources): A) Colored Glass Filters on traditional TQH sphere sources
Standard sphere source with colored glass filter
Results
Light gray line: what the instrument measured
Dots: what it should have gotten
Solid line: example of stray light correction algorithm (developed using laser characterizations)

## Slide 10
Evaluate for Presence of Effect (2)
Novel sources with well defined and distinct spectral shapes (compared to tungsten quartz halogen sources): B) Solid state (LED) sources
A wide variety of LED choices exist; measurement of the “tails” is a good test of a spectroradiometer
The actual spectral radiances would be determined by a double grating or prism grating system at NIST

## Slide 11
Ideas & Suggestions (4)
The wavelength calibration should be evaluated frequently. It can depend on temperature, time, etc. Data from Milton Hom, NASA / GSFC

## Slide 12
Recommendations to IVOS
NIST is available to help with laboratory component, including planning and execution, of any CEOS/IVOS sponsored field comparison activity. This activity would be coordinated with our collaborators at NASA, IPO, and NOAA and the level of our involvement would depend on the interest and enthusiasm of CEOS / IVOS as well as these agencies.